| Sr.No | Register 40003 | Register 40004 | Decimal | Value | Difference |
| --- | --- | --- | --- | --- | --- |
| 1 | 007FH | 0B18H | 8325912 | 9.925 | 0.25% |
| 2 | 007FH | 0EC4H | 8326852 | 9.926 | 0.26% |
| 3 | 007FH | 0BBBH | 8326075 | 9.925 | 0.25% |
| 4 | 007FH | 328CH | 8336012 | 9.937 | 0.37% |
| 5 | 007FH | 348EH | 8336526 | 9.937 | 0.37% |
| 6 | 007FH | 33FFH | 8336383 | 9.937 | 0.37% |
| 7 | 007FH | 3C36H | 8338486 | 9.940 | 0.41% |
| 8 | 007FH | 3623H | 8336931 | 9.938 | 0.38% |
| 9 | 007FH | 39BCH | 8337852 | 9.939 | 0.39% |
| 10 | 007FH | 3E84H | 8339076 | 9.940 | 0.41% |
| 11 | 007FH | 37DAH | 8337370 | 9.938 | 0.38% |
| 12 | 007FH | 3C82H | 8338562 | 9.940 | 0.41% |
| 13 | 007FH | 390EH | 8337678 | 9.939 | 0.39% |
| 14 | 007FH | 3DABH | 8338859 | 9.940 | 0.41% |
| 15 | 007FH | 3D51H | 8338769 | 9.940 | 0.41% |
| 16 | 007FH | 3955H | 8337749 | 9.939 | 0.41% |
| 17 | 007FH | 4001H | 8339457 | 9.941 | 0.42% |
| 18 | 007FH | 40F0H | 8339696 | 9.941 | 0.42% |

Difference formula: Difference=(test value-standard value)/standard value*100%

Test value formula: 40003 low 8-bit is the high 8-bit of this channel, 40004 is the low 16-bit value of this channel; convert 24-bit data to Decimal data, then divided by 8388607, finally multiply by span, that is the voltage.

Online upgrade steps:

Step 1: Open UT-IAP V1.0 software, choose the serial port, then open the serial port

Step 2: Load the “bin” file, then start upgrade

Step 3: Cut off the power of the device, reboot, the device is entering the upgrade status.

Step 4: After upgrade, the device restart, upgrade is finished

After upgrade, it adopts standard MODBUS protocol

| Register address | Description |
| --- | --- |
| 30001 | Channel 0 high 16-bit |
| 30002 | Channel 0 low 16-bit |
| 30003 | Channel 1 high 16-bit |
| 30004 | Channel 1 low 16-bit |
| 30005 | Channel 2 high 16-bit |
| 30006 | Channel 2 low 16-bit |
| 30007 | Channel 3 high 16-bit |
| 30008 | Channel 3 low 16-bit |
| 30009 | Channel 4 high 16-bit |
| 30010 | Channel 4 low 16-bit |
| 30011 | Channel 5 high 16-bit |
| 30012 | Channel 5 low 16-bit |
| 30013 | Channel 6 high 16-bit |
| 30014 | Channel 6 low 16-bit |
| 30015 | Channel 7 high 16-bit |
| 30016 | Channel 7 low 16-bit |

| Register address | Description |
| --- | --- |
| 40001 | Channel 0 range span |
| 40002 | Channel 1 range span |
| 40003 | Channel 2 range span |
| 40004 | Channel 3 range span |
| 40005 | Channel 4 range span |
| 40006 | Channel 5 range span |
| 40007 | Channel 6 range span |
| 40008 | Channel 7 range span |

Span description：

1：±10V

2：±5V

3：±20mA

4：±1V

5：±500mV

6：±150mV

| Register address | Description |
| --- | --- |
| 45001 | Device address |
| 45002 | Baudrate |

Remarks:

- Device address range 0~247

- Baudrate

0: 1200

1: 2400

2: 4800

3: 9600

4: 19200

5: 38400

6: 57600

7: 115200